## Slide 1: Results of the Microwave Transmission Measurements
MD Runs 22. + 31.10., 02. + 05. 11., 05. 12. 2012 F. Caspers, S. Federmann
07. 02. 2013

## Slide 2: Overview
Measurement Setup
General Remarks
Objectives
Results
Conclusion

## Slide 3: Measurement Set-Up
uncoated
uncoated
Surface at BA5

## Slide 4: Objectives
This talk summarizes the results of eight MD runs (March to December 2012)
The electron cloud induced phase modulation signal was studied for high intensity beams
The electromagnetic ‘cleanliness’ of the new setup was tested
The functionality of the filter in the receiver path of the setup was tested

## Slide 5: Results
3 batches with 25 ns spacing
3 batches with 50 ns spacing
Calibration signal @ 42kHz
Modulation signal @ 43.35kHz
PM Demodulation
PM Demodulation

## Slide 6: Results
PM to AM conversion was detected
Determined to be 8% -> negligible for our ecloud singal (height: roughly 5 dB -> loss of 0.4 dB)
No beam in SPS
SG connected directly to VSA
PM Demodulation
AM Demodulation
PM Calibration signal
PM Calibration signal in AM modulation
PM Calibration signal

## Slide 7: Results
Filter is absolutely crucial for our setup
Without it, lots of fake lines due to intermodulation are present
With filter in receiver path
Without filter in receiver path
Strong fake signal due to intermodulation
PM Demodulation
AM Demodulation

## Slide 8: Results
Reduction of electron cloud signal from 7 to 3 - 5 dB above noise floor
Possible scrubbing?
2 effects visible:
  Strong signal at injection (not fully understood)
  Virtually no change in steady state but stronger fluctuations
25 ns beam, 3 batches March 2012
25 ns beam, 3 batches November 2012
+3-5dB
+7dB

## Slide 9: Results
Injection effect also observed for 50 ns spacing
In March 2012, no injection signals are present whereas in November 2012 they are
50 ns beam March 2012
50 ns beam November 2012

## Slide 10: Conclusions
Experimental setup works properly -> former findings could be confirmed (25 ns electron cloud signal present, 50 ns no signal visible)
Cleanliness of the experimental setup is due to use of filter -> confirmed for this setup
Potential scrubbing effect seen. “Same” beam (25 ns) caused a higher net signal height in March 2012 than in November 2012

## Slide 11: Thank you for your attention!

## Slide 12: Additional Slides

## Slide 13: Theoretical Basis
Principle: Measurement of the induced phase shift of a microwave going through a plasma filled waveguide
Phase shift is proportional to the electron cloud density:
Assumption: we have a complete transmission through the electron cloud (no resonances)

## Slide 14: Screenshot of the VSA