## Slide 1: The SPS Transverse Damper Status 2015 / Plans 2016
G. Kotzian for the SPS Damper Team
BE-RF-FB, 03. December 2015 LIU-SPS BD WG
http://paf-spsu.web.cern.ch/paf-spsu/
Acknowledgements: BE-RF-FB,
BE-RF-CS,
BE-RF-PM,
BE-OP,
BE-ABP

## Slide 2: Outline
Objective:
Have a look at the upgraded SPS Damper, illustrate few components which have been evaluated in 2015, and how the damper could be used e.g. for MDs in 2016.
Structure:
Design of the upgraded SPS Damper
Status 2015
Outlook for 2016 (briefly)

## Slide 3: Transverse Damper in General
Kicker
The transverse damper is a feedback system: it measures the bunch-by-bunch oscillations and damps them by fast electrostatic kickers.
Analog
Front
End
SUM
Signal
Processing
BPM
Power Amp
ADC
DAC
5ns Bunches
Active closed loop feedback
transverse
position
pre-processed
sampled
position
calculated
correction data
correction
signal
drive signal
BBQ

## Slide 4
PU1
BEAM
BPM
Proc.
PU2
BPM
Proc.
BPM
Proc.
BPM
Proc.
Electrostatic PU processing
BW ~100MHz
Stripline pickup
processing
Damper Signal Selector
BeamPos1
BeamPos2
DSPU1
DSPU2
ADC1
Baseband Frontend
(40 MHz)
ADC2
DAC1
DAC2
BeamPos1
BeamPos2
DSPU1
DSPU2
ADC1
RF Frontend
200 MHz FSK
ADC2
DAC1
DAC2
BeamPos1
BeamPos2
DSPU1
DSPU2
ADC1
Baseband Frontend
ADC2
DAC1
DAC2
BeamPos1
BeamPos2
DSPU1
DSPU2
ADC1
RF Frontend
200 MHz
ADC2
DAC1
DAC2
from BBQ
Switch M1
Switch M2
pLHC
pLHC
pFT
pFT
Ions
Ions
Scrub
Scrub
from VNA
from VNA
SURFACE BA2
TUNNEL LSS2
pLHC,
Scrub
Ions
pFT
KICKER1
PLC1
PLC2
200 W
Driver
200 W
Driver
Final
Stage
Final
Stage
Damper Power System
KICKER2
SPS Damper Loops
pLHC
SPS Damper Loops
pFT
SPS Damper Loops
Ions
SPS Damper Loops
Scrub
cfv-ba2-alltrdamp[hv]1
cfv-ba2-alltrdamp[hv]2
LLRF MODULES:
6 commissioned
2 Experimental
2 Ions (to do)
POWER:
Excellent availability during 2015!

## Slide 5: 6x Beam Position Monitors installed in LS1
BPCR.221
all BPMs re-cabled with 7/8-inch smooth wall coaxial cables during LS1; [length ~250..700m]
BPH/BPV: electronics and amplifiers in the tunnel tested and verified; "hot-spares" readily checked and kept in the tunnel
BDH / BDV kickers unchanged
RF Faraday cage
SPS Transverse FB
2 BPCR (H/V) for LHC type beams
(couplers maximum ZT @ 200 MHz)
2 BPH electrostatic PU (pFT)
2 BPV electrostatic PU (pFT)
BPH/BPV electronics (in pits right below the BPMS) are extremely robust.
Very reliable also during operation 2015.

## Slide 6: BPCR Stripline Pickups
(B) BPCR.221.V shows “echo”
after beam passage.
Lasts for ~1.77 us
Appears to not depend on transverse position
“ECHO”
≈1.77 µs
CH2: SUM
CH4: DELTA
(A) BPCR.214 at FT, 12 bunches.
Raw DELTA signal reveals frequencies > 1 GHz being picked up.
discovered 2014/11/08 (origin unclear, tunnel inspection foreseen)
BPCR.221
Build-up during beam passage
Decay ~ 100 ns
CH1: SUM-H
CH2: DELTA-H
CH3: SUM-V
CH4: DELTA-V

## Slide 7: Analog Frontend Control / External Control
Two analog frontend types:
  RF frontend, used for pLHC and Ions, LO = 200 MHz, I/Q components
  Baseband frontend (BB), used for pFT and Scrub, only I-component
External control:
  Input termination / attenuation / amplifier
  Hardware missing for Ions (foreseen 2016)
Input gain control: RF frontend
  RF: switching attenuators + amp, delay matched Large dynamic range: -39 dB to + 20 dB = 59 dB
Input gain control: BB frontend
  BB: variable gain amplifier (VGA)
  Large dynamic range: -27.4 dB to + 26 dB = 53.4 dB
  Signal levels between VGA and ADC not inline
  Patching of signal levels required  saturation+noise
LO 200 MHz
120 MHz
120 MHz
RF frontend schematic

## Slide 8: 200 MHz Comb Filter
5 ns
THRU line of stripline coupler picks up “ECHOs”
Response not flat
 Baseband pulse distorted
Affecting SCRUB Frontend
Spoiling pLHC response (200 MHz)
25 ns
Undershoot (cross-talk)

## Slide 9: Beam Position Block (BeamPos)
2 BeamPos Blocks per damper loops module = one per pick-up input
I/Q inputs of SUM and DELTA signals
  in ADC sampling clock domain, hence subject to delay adjustment
CoRe – Conditioning and Resampling
  I/Q phase rotation
  Down sampling: 120 MSPS  40 MSPS (25ns bunch sync)
  Combining 3 samples for one position signal
   sampling delay problematic already at 200 MHz
Normalization
  position calculation (in-phase)
  “quadrature position” indicating head-tail activity
  bypass mode in case of DELTA processing (pFT)

## Slide 10: Analog I/Q inputs of SUM and DELTA
Skew between I and Q signals
Time lag beween I/Q up to 3.5 ns
Depending on phase of LO phase
SAMPLING POINTS
Bunch n-1
Bunch n
Bunch n+1

## Slide 11: Analog I/Q inputs of SUM and DELTA
Skew between I and Q signals
Time lag beween I/Q up to 3.5 ns
Depending on phase of LO phase

## Slide 12: Transverse Position
Slope  bunches towards the end of a batch appear on ‘different orbit’
pLHC module (200 MHz analog frontend)
batch of 72 bunches, nominal intensity
Aquired after decay of injection oscillations  bare orbit = transverse offset (DC)
Potential solution: use direct sampling, also for 200 MHz input – first tests with beam affirmative …
Implications on damping?

## Slide 13: DSPU – Digital Signal Processing Unit
Functions (from RFFG)
Control signals (FESA):
Excitation input (DDS)

## Slide 14: DSPU – Digital Signal Processing Unit
2 DSPUs per damper; one per module; two inputs
Signal processing
  phase equalizer FIR (up to 64 taps), linearize phase of power system ≈ low pass 1st order, cutoff freq. (-3dB) @ 4.5 MHz operated at 40 MSPS  better resolution at lower frequencies
  gain equalizer (up to 64 taps), digital low pass shaping output signal operated at 120 MSPS
Recently …
  Flat phase equalizer FIR and
  Flat gain equalizer FIR
phase equalizer FIR, comp. 1st order 4.5 MHz
digital low pass,
fc > 20 MHz

## Slide 15: 2015-11-11
http://elogbook.cern.ch/eLogbook/event_viewer.jsp?eventId=2039845
“Dampers are happily running at maximum bandwidth […]. Current modules: DamperLoop_H_pLHC and DamperLoop_V_pLHC”
  For discussion:
  Flat phase equalizer FIR and
  Flat gain equalizer FIR
NB: tunes are at
Qh = 0.13
Qv = 0.18

## Slide 16: Semi-Auto Parameter Extraction
basis is the classical method with VNA(s)
  automation with scripts
  automatic data collection
Measure OL 1 MHz to 20 MHz
  (here: Meas39, BPCR.221.V  V2)
  20 sweeps x SC period … 15-30 min
Post-processing allows parameters to be extracted, e.g.
  Phase advance PU  Kicker
  Feedback COARSE + FINE delay
  Tune
  Loop phase response
PRELIMINARY (Meas39 - 2015/11/04)
Open loop response (polar):

## Slide 17: Fractional Tune
PRELIMINARY (Meas39 - 2015/11/04)
Open loop response (mag):
Open loop response (mag):
Overlay of 20 measurements: from 1 MHz to 20 MHz
Feature extraction: blue = q green = (1-q)

## Slide 18: Loop Phase Response
PRELIMINARY
NB:
The Power system is considered to have a
1st order low-pass characteristic,
with cutoff-frequency 4.5 MHz (=45 Grad)
… for discussion
4.5 MHz

## Slide 19: USER INTERFACE
GUI in JAVA
  talks to LSA (set parameters) and FESA (read back)
  loads settings to LSA for 1 active damper (sets others to “safe settings”)
  display status and monitoring
  trimming of feedback parameters (in LSA, with history)
  store/recall reference settings
  import/export damper settings to file (consistent snapshot, e.g. for/during MDs)
“RESTORE TO
DEFAULT SETTINGS”

## Slide 20: Drive Sequence
Copy REFERNCE settings to active modules
Load SAVE settings to inactive modules

## Slide 21: Status 2015
  Signal processing Modules, aka. DamperLoops:
    6 commissioned
    2 Experimental
    2 Ions modules, in preparation for 2016
  Damper Power system
    Excellent availability during 2015
    Special thanks to colleagues from RF-PM!
  Pickup Signal processing
    Electronics in the tunnel extremely robust & very reliable
    Performance of analog signal processing chain under evaluation
  Observation Box
    Prototyping: server installed, signalling already running with test crate
  First MD studies on damper performance evaluation
  … running also for Ions (using pLHC temporarily)

## Slide 22: Plans for 2016
  Evaluating concepts for improvements of the damping performance
    different implementations for phase shift filters
    adaptive feedback control
    test a novel damping algorithm
    bunch extrapolation (single bunches, ions), beam gap interpolation (in between batches)
  Develop methods to facilitate loop diagnostics
    loop information indispensable during commissioning and for setting-up
    easy-to-use repeatable measurements: allows re-qualification of settings/parameters  handy e.g. for cycle generation
    review potential for per-bunch head/tail activity monitoring (test with wideband feedback)
  exploitation of new features
    real bunch-by-bunch damper (in view of crab cavity tests)
    adding a DDS allows for built-in-network analysis
    controlled emittance blow-up, beam cleaning, and bunch shaping using transverse excitation signals

## Slide 23: Feedback performance evaluation with different phase shift filters
Objective: study different implementations of phase shift filters
Goal: make feedback phase-insensitive to tune variations
Phase compensation for sidebands
Notch, 1T delay, phase rotation

## Slide 24: Real bunch-by-bunch damper based on feedback gain equalization
Objective: shaping the time-domain response to generate individual kicks per-bunch, requires feedback gain equalization/gain levelling
LHC case  see W. Hofle, “PERFORMANCE OF THE LHC TRANSVERSE DAMPER WITH BUNCH TRAINS,” WEPME043 - IPAC’13.
Useful for crab cavity tests
Individual bunch blow-up (in coast)
Fully linearized system
LHC ADT

## Slide 25: THANK YOU
Questions?